| Met One Instruments, Inc. 1600 NW Washington Blvd, Grants Pass, OR TEL (541) 471-7111 Fax (541) 471-7116 metone@metone.com |
| --- |
| Title: BAM-1020 Audit Sheet and Check List Sample |
| Document Number: | Revision: A |
| Originator Approval: | Checker Approval: | Document Control: |

| Revisions |
| --- |
| Rev | Description | Date | By |
| 1 | 1st Draft. Derived from the old BAM-1020 Audit Sheet. | 12/27/2005 | DAH |
| A | Initial Release. | 3/20/2006 | DAH |
| B | Revised for Rev F manual insertion. | 3-2007 | DAH |
| C | Revised for PM2.5 default settings and three point flow calibrations. | 4-7-2009 | DAH |

BAM-1020 Audit Sheet

| Model: | BAM-1020 |  | Serial Number: |  |
| --- | --- | --- | --- | --- |
| Audit Date: |  |  | Audited By: |  |

| Flow Audits |
| --- |
| Flow Reference Standard Used: | Model: |  | Serial No: |  | Calibration Date: |  |
| Temperature Standard Used: | Model: |  | Serial No: |  | Calibration Date: |  |
| Barometric Pressure Standard Used: | Model: |  | Serial No: |  | Calibration Date: |  |
| Leak Check Value: | as found: | slpm |  |  | as left: | slpm |  |  |  |
|  |  | BAM | Ref. Std. |  |  | BAM | Ref. Std. |  |  |
| Ambient Temperature: | as found: | C | C |  | as left: | C | C |  |  |
| Barometric Pressure: | as found: | mmHg | mmHg |  | as left: | mmHg | mmHg |  |  |
| Flow Rate 15.0 lpm: | as found: | lpm | lpm |  | as left: | lpm | lpm |  |  |
| Flow Rate 18.4 lpm: | as found: | lpm | lpm |  | as left: | lpm | lpm |  |  |
| Flow Rate 16.7 lpm: | as found: | lpm | lpm |  | as left: | lpm | lpm |  |  |

| Mechanical Checks |
| --- |
| Pump muffler unclogged: | as found |  | as left |  | PM10 particle trap clean: | as found |  | as left |  |  |  |
| Sample nozzle clean: | as found |  | as left |  | PM10 drip jar empty: | as found |  | as left |  |  |  |
| Tape support vane clean: | as found |  | as left |  | PM10 bug screen clear: | as found |  | as left |  |  |  |
| Capstan shaft clean: | as found |  | as left |  | PM2.5 cyclone particle trap clean: | as found |  | as left |  |  |  |
| Rubber pinch rollers clean: | as found |  | as left |  | Inlet tube water-tight seal OK: | as found |  | as left |  |  |  |
| Chassis ground wire installed: | as found |  | as left |  | Inlet tube perpendicular to BAM: | as found |  | as left |  |  |  |

| Analog Voltage Output Audit | N/A |  |
| --- | --- | --- |
| DAC Test Screen | BAM Voltage Output | Logger Voltage Input |
| 0.000 Volts | Volts | Volts |
| 0.500 Volts | Volts | Volts |
| 1.000 Volts | Volts | Volts |
| Span Membrane Check |  | BAM Clock |
| LAST m: | mg/cm2 |  | BAM Date: |  |
| ABS: | mg/cm2 |  | BAM Time: |  |
| Difference: | mg/cm2 |  | Actual Time: |  |
| Difference: | % |  | Set? (Y/N) |  |

| SETUP > SAMPLE Settings | SETUP > CALIBRATE Settings | Other Settings |
| --- | --- | --- |
| Parameter | Expected | Found | Parameter | Expected | Found | Parameter | Expected | Found |
| RS232 | 9600 |  | FLOW RATE | 16.7 lpm* |  | e1 | -0.015 mg |  |
| BAM SAMPLE | 42 min* |  | FLOW TYPE | ACTUAL* |  | Cycle Mode | STANDARD |  |
| STATION # |  |  | CONC TYPE | ACTUAL* |  | RH Control | YES* |  |
| MET SAMPLE | 60 min |  | Cv |  |  | RH Setpoint | 35%* |  |
| RANGE | 1.000 mg |  | Qo |  |  | Datalog RH | YES |  |
| OFFSET | -0.015 mg |  | ABS |  |  | Delta-T Control | NO* |  |
| CONC UNITS | mg/m3 |  |  sw |  |  | Delta-T Setpoint | 99 |  |
| COUNT TIME | 8 min* |  | K |  |  | Datalog Delta-T | NO |  |
|  |  |  | BKGD |  |  |  |  |  |
|  |  |  | STD TEMP | 25 C |  |  |  |  |
|  |  |  | HEATER | AUTO* |  |  |  |  |

* Required PM2.5 FEM settings. See manual.

| Last 8 Errors in the BAM-1020 Error Log |
| --- |
| Error | Date | Time |  | Error | Date | Time |
| 1 |  |  |  | 5 |  |  |
| 2 |  |  |  | 6 |  |  |
| 3 |  |  |  | 7 |  |  |
| 4 |  |  |  | 8 |  |  |

Notes:

# Text boxes

About this Document

This Audit sheet is a recommended list of system checks for the Met One Instruments BAM-1020 unit. Some of the line items in this form may not be applicable to certain units due to optional hardware. Certain units may need additional checks not listed on this sheet.

To use this form for system audits, fill out the “as found” items without making any changes to the unit. To use this form for a test and calibration record, fill out the “as found” and “as left” items, performing any necessary maintenance and calibrations required.

This document may be modified, revised, translated, or incorporated into SOPs by the end user as needed.

- Flow Audits:		Perform an as-found leak check, then clean the nozzle and vane and perform an as-left leak check. Perform a flow audit or calibration, recording the values before and after making any changes.

- Mechanical Audits:	Inspect each of the listed items and perform any maintenance necessary. Check each box as the item is completed.

- Analog Voltage Audit:	Enter the TEST>DAC screen and force the voltage output to each level, then measure the voltage at the back of the BAM and at the analog data logger to verify agreement. These checks are not required if no external analog data logger is used.

- Membrane Audit:	Record the last span membrane measurement (LAST m) from the OPERATE>NORMAL screen and compare it to the ABS value. Calculate and record the difference.

- Setup Values:		Enter the SETUP menus and verify each setting value has not been changed. Items marked with an asterisk are critical PM2.5 FEM settings. Enter the factory calibration settings from the unit’s calibration certificate. The rest of the expected values should be recorded when the instrument is initially set up, since some of the settings will differ for various applications.